Phosphoimager Instructions

Exposing

- Turn on surge protector

- Put image sheet on the light box (exposing surface against light) for 15 minutes

- Lay dried gel in cassette

- Cover gel with the image sheet and close cassette

- Turn surge protector off

Reading

- Turn on surge protector

- Turn on computer (watch for on-screen delays)

- Open door on imager (on the front surface)

- Pull drawer out

- Open cassette and remove image sheet, noting the coordinates to be scanned

- Lay image sheet face up in the drawer

- Close the drawer and the door

- Enter user name and password into the computer (found next to the computer)

- Click “Quantity One” icon on the desktop

- Under the “File” tab, click “Personal Fx”

- Enter the scanning coordinates in the boxes labeled “Top”, “Bottom”, “Left”, and “right”

- Click “Acquire”

- When scanning is finished, click the “File” tab, click on “Save As” and save the image under My document, your lab’s file, your file name

- Put image sheet on the light box, (exposing surface against light) for 15 minutes

- Click “Zoom Box” to get the cropped image

- Click “Transform” to adjust the image

- Click “Volume Rect tool” and draw a box around the most pronounced band (note which band you chose and the order of bands selected in the next steps)

- Push “Control” and click Left Mouse to make a copy of the box made in step 17

- Move the box over the next band to be quantified

- Repeat steps 18 and 19 until all bands are surrounded by a box

- Copy one more box to surround a blank gel area

- Label each box as follows: Double click on the box

Click next to either “Unknown” for a sample, or

“Background” for the box surrounding the blank area

Type the sample name (or background) in the Name

slot

Click OK

- Click “Volume Analysis Report” to display results

- Click “O.K.”

- Click the icon that looks like a table to export and save the results

- Open the results and adjust format the way you like to see the data

- Also save to a disk and/or send as an email

- Turn off the computer

- Turn off the surge protector

- Put the image sheet in the cassette